REQUEST FOR PROPOSAL

HEMODIALYSIS MACHINE SUPPLY

AND HOSPITAL BEDS ELECTRIC

REFERENCE NO: SO/2026/ 000010033

Date: 13ST January 2026

- ABOUT QATAR CHARITY:

Qatar Charity is one of the leading non-profit organizations in the Arabian countries and the world. It was established in 1992 for the purpose of developing Qatari communities and the surrounding destitute communities regardless of their race, gender, and cultural differences.

Qatar Charity is active in several sectors significantly in sustainable development, fighting against poverty, and relief for the emergency afflicted situations. Qatar charity enjoys the consultative status with the Organization of Islamic Cooperation as it is a founding member of the Arab Network for Development in Cairo since 1999.

QC has been working in Somalia since 2007. It’s mainly working in the following thematic areas: Education, Health, Shelter, Livelihood, Emergence, social care, WASH and construction of public facilities and CPI. QC currently conducting health strengthen projects in Somalia by working with minister of health at federal and state level. The aim of health program is to ensure that all people get continued access to health services.  In this Request for proposal, QC is going to Equip and furnish the new soon-to-be completed De-Martino Hospital’s emergency and trauma center funded and constructed by Qatar Charity.

- PURPOSE OF THIS REQUEST FOR PROPOSAL:

Qatar charity is seeking qualified vendors/Suppliers whose specialty in medical equipment in Somalia and Abroad. All potential Vendors/Suppliers are requested to submit their required documents according to the Conditions as detailed further made herein below tender descriptions.

- LOCATIONS:

Mogadishu, and Hargeisa.

- SCOPE OF THE WORK:

The scope of this RFP includes the procurement, purchase, delivery, and install of medical equipment’s.

- Medical equipment and Devices:

- The Bill of Quantities (BOQs) provided by the Office Qatar charity contains the minimum specification that is required for the item. As per the terms of the agreement, the contractor/vendor is required to provide comprehensive details of the item being quoted for, including the brochure and specifications. This information should be included in the quotation provided by the contractor/vendor.

- for any medical equipment or devices that quoted, the price should include both installation and maintenance services for a period of no less than 12 months. To ensure that the equipment will be properly set up and maintained for at least half a year after purchase. Also it’s required from the vendor to be able to supply spare parts of the medical devices and equipment for a period of 6 years.

- Training the hospital's medical staff on operating the devices.

- DOCUMENTS REQUIRED TO SUBMIT SEPERATELY

- Updated Business profile.

- Valid License/ Registration from the concerned authorities in Federal/state level.

- Initial Guarantee cheque/letter $5,000 (Ten thousand US Dollars)

- Tax Clearance certificate (TCC):  a valid tax compliance certificate from Ministry of Finance.

- Completed (Filled and Stamped) Tender Document (This Document)

- ID or Passport of the Owners.

- Bank certificate should include the following details, Account Name, Account Number, Iban Number and Bank signatories.

- Expression of interest letter

- Shall submit at least three previous similar supply and delivery of medical equipment related contracts implemented with and/or for government institutions, International non-governmental organizations (INGOs), UN agencies or private institutions.

- Key personal of the company including biomedical Engineer’s CV

- Any information deemed to be important.

- GENERAL TERMS AND CONDITIONS

- Only Suppliers that are legally registered in Somalia can apply for this Request for proposal.

- Only the original RFP form received from Qatar Charity - Somalia Office (QC) should be used for submitting the offer. However, in case of insufficient space, additional sheets may be attached retaining the same column heading as the original Tender Form.

- All the columns provided in the RFP documents should be filled in. These columns should not be left blank. In the absence of any information, please write N/A.

- Overwriting, erasing, or using correction fluid are not allowed.

- Implementation time should always be mentioned in number of Calendar Days.

- The validity of the offer shall be minimum 180 days from the closing date of submission.

- QC has the right to revise the Supply Details as applicable.

- The Bidder shall bear all costs associated with the preparation and submission of his Bid, and the Qatar Charity shall not be responsible or liable for those costs, regardless of the conduct or outcome of the bidding process.

- QC has the right to cancel any item(s) from the Request for proposal or to cancel the whole Tender. QC is under no obligation to accept the lowest price or any offer, offer always is separate-able, and QC has the right to take part of it.

- The Offer shall include full details and specifications of the offered items.

- For DUE DILIGENCE CHECKS purposes, QC has the right to ask Bidders to provide any extra information. Before or after the final Selection, Bidders may possibly be required to attend a meeting with the Procurement Committee in QC at no extra cost to QC.

- QC has the right to request submission of Company Registration Documents and Passport Bio-Page of key personal of Supplier as issued by the Government.

- Please note that once a Tender process completed, QC will ONLY communicate with the SELECTED BIDDER going forward.

- Unsuccessful Bidders reserve the right to know the results by inquiring from QC if interested. All Suppliers/Vendors are welcomed to ask questions for the reason(s) why they did not win. Bidder information will NOT be shared with a Third Party.

- BASIC INFORMATION ABOUT THE VENDOR/ HEALTH SERVICE PROVIDOR

| A. HEALTH SERVICE PROVIDOR DETAILS: |  |  |
| --- | --- | --- |
|  | Name of Company |  |
|  | Address |  |
|  | Telephone No. |  |
|  | Mobile No |  |
|  | Fax No |  |
|  | Type of Business |  |
|  | Number of Years in business |  |
|  | Contact Person |  |
|  | Mobile No. |  |
|  | E-mail |  |
|  | Company Owner |  |
|  | Company Manager |  |
| B. LICENCE DETAILS |  |  |
|  | Registration No. |  |
|  | Expiry Date |  |
| C. BANK DETAILS |  |  |
| 1 | Name Of Bank (1) |  |
|  | Account No. (1) |  |
| 2 | Name Of Bank (2) |  |
|  | Account No. (2) |  |
| D. SIGNATURE&STAMP OF THE COMPANY: |  |  |
| Name: Stamp Title: Signature |  |  |

- SELECTION CRITERIA:

- Eligibility Evaluation:

The items mentioned below are evaluated according to the documents attached to the company in the technical offer, and the evaluation is based on their availability and the validity of the document only.

| # | Item | Yes/No |
| --- | --- | --- |
|  | Valid License/ Registration from the concerned authorities at Federal/state level. |  |
|  | Tax Clearance certificate |  |
|  | Preliminary Guarantee cheque ($5,000) |  |

- If the supplier gets a “NO” from this evaluation will be disqualified automatically to proceed to the next step.

| Registration Documents | Tender Conditions + Legal Documents | Company Profile (Company Overview – Areas of Operation – Establishment – Executed Projects) | Company's general Experience in the field | Experience in the field with a list of projects executed over the past four years | Inspection of the submitted samples and evaluation of their quality and compliance with specifications |
| --- | --- | --- | --- | --- | --- |
| essential condition | 10.0% | 10.0% | 20.0% | 35.0% | 25.0% |

- Technical Evaluation:

| # | Item | Marks |
| --- | --- | --- |
|  | Tender Conditions + Legal Documents | 10 |
|  | Company Profile (Company Overview – Areas of Operation – Establishment – Executed Projects) | 10 |
|  | Company's general Experience in the field | 20 |
|  | Experience in the field with a list of projects executed over the past four years | 35 |
|  | Inspection of the submitted samples and evaluation of their quality and compliance with specifications | 25 |
| TOTAL |  | 100% |

- Only Bidders that obtain minimum 70% of the Evaluation Based on the Technical criteria will be considered for Financial Evaluation.

- Financial

| # | Item |  | Description | Quantity | Location | Total |
| --- | --- | --- | --- | --- | --- | --- |
| 1 | HEMODIALYSIS MACHINE | DIALYSIS DESCRIPTION: - The machine should have to have a facility for conventional and High flux dialysis - The machine should have two bacterial filters (Pyrogen filters) i.e. one at water inlet site and the other one before water going to dialyser - The battery should have to be used as a back – up for 20 – 30 minutes and should have to run the machine completely with heater supply - The machine should have to have Na, Bicarbonate and UF profiling - The machine should have dialysate temperatures that should have to be selectable between 35 to 39 Degree Centigrade - The machine should have variable conductivity setting ranging between 12 to 15 - The machine should have to have variable dialysate flow ranging between 200- 800 ml/minutes and should have increasing facility to step up by 20 ml - The machine should have the facility to show trends curve of all parameter for 15 – 20 minutes; the machine should have Heparin pump with syringe sizes 20 to 30 ml with pump flow rate from 1-10 ml/hour (with 0.1 ml increments) - The machine should have to have stroke pressure that is going to be operated with short term single needle dialysis. - The machine should have to have ultrafiltration that should function at a rate of 0.1 to 2.5 liters/hr. The in and out fluid circuit must have to be separated to avoid contamination in the event of membrane rupture. - The treatment parameter of the machine should have to be displayed both graphically and digitally - The machine should have to have integrated heat and chemical disinfection facility / system with one standard time - The machine should have to have disinfection program with day and night arrangement with weekly schedule - The machine should have an accurate feedback control conductivity mixing technique - The machine should have to have drain facility / system - The machine should have to have an accurate UF control by flow and by volume control measurement techniques - The machine should have to have automatic self - test facility - The machine should have auto ON/ OFF facility / control system - The machine should have to have touch button screen and large colour TFT Screen - The machine should have to have automatic diagnosis of malfunctioning with on line ability to show the faults with trouble (Technical Service Mode) - The machine should have to be connected to computer to feed all data and troubleshoot whenever any problem - The machine should have to have a blood pump rate from 20-500 ml/min adaptable to standard A-V bloodlines of 140 - The machine should have to have the ability to monitor pulse rate and NIBP - The machine should have to have audio visual alarms on limit violation of conductivity, blood leak, air leak, trans - membrane pressure alarms, dialysis temperature alarm, dialysis can empty alarm, end of disinfection alarm, bypass alarm and blood pump stop alarm - The machine should have to have an alarm for reverse Ultrafiltration and also be able to do sequential dialysis - The machine should have to have online in built NIBP recording - The machine should have to have online clearance monitoring built in device for measurement and monitoring of effective urea clearance and dialysis dose (KT/V) - The machine should have to have power input to be 220-240VAC, 50Hz fitted with standard Ethiopian plug. - The machine should have to be FDA, CE, UL or BIS approved product - The manufacturer/Supplier of the machine should have to have ISO certification for quality standards. - The machine should have to comply with IEC 60601-2-16 SAFETY requirements of medical electric equipment part2-particular requirements for the safety of Hemodialysis equipment Dialysis Chair. |  | 22 |  |  |
| 2 | HOSPITAL BEDS ELECTRIC | Four split side rails, drop down. Detachable head/footboard with easy lock Central brake system with a wheel diameter of 125mm Bumper caster on four corners Auto-counter to prevent the patient from sliding onto the foot end ABS articulated soft joint for the platform With an angle indicator embedded in the railing for back and Trendelenburg adjustment Manual CPR on both sides. Overall size 2200*1085*450mm Recommended mattress size 2015*875mm. Height adjustable: 350-700mm Trendelenburg/reverse Trendelenburg: 0-12° Loading capacity: 225kg Accessories: Head/Foot Board 2pcs Side Rails 4pcs Bed Frame 1Set Bed Boards 5pcs Casters 4pcs IV Pole with 4holders 1 Pcs Drainage Bag Holders 2pcs Overbed Table Specification Tabletop Medium density fireboard Table stand Steel Dimension L600*W400*H600-850mm, Tabletop dimension L600*W400mm, Thickness 15mm Features, Height adjustable: 600-850mm, The table top can be rotated 360° and tilted back and forth 45°. With castors, easily moving, Supplying the biggest convenience to patients. Bedside Table Hospital Furniture, Bedside Cabinet Table * Dimension 2: L450×W420*170 ×H900mm * Material carbon steel spraying. ABS plate. |  | 22 |  |  |
| 3 | Blood pressure, digital | Potential biders are required to discribe the detials of the machine you are quoting to in terms of: 1- Manufacturer? 2- Model? 3- Power Inputs? 4- Year of production 5- Screen interface? 6- Control and mointor systems? 7- Attached accessories |  | 2 |  |  |
| 4 | Waste bin with lid | Polished Stainless Steel with a capacity 5Litr |  | 2 |  |  |
| 5 | Drug cabinet | Medicine display cabinet with blind doors and glass doors. Structure made of melamine. Internal shelves in melamine with ABS profile adjustable in height. Glass doors on perimeter frame in aluminum with central mirror in transparent tempered glass thickness 4mm Blind doors made of molded and shaped MDF. Dimension: Not Less than WxDxH 90x40x200cm |  | 2 |  |  |
| 6 | Nurse Desk |  |  | 2 |  |  |
| 7 | Chair |  |  | 2 |  |  |
| 8 | File Cabinet |  |  | 2 |  |  |
| 9 | Drug cabinet |  |  | 2 |  |  |
| Total |  |  |  |  |  |  |

- APPLICATION PROCESS

Potential Suppliers having interest the work described above shall submit their Interest and supporting documents.

The Expression of Interest along with accompanying documents must be received by QC no later than Thursday 22nd January 2026 @ 2:00pm.

You are requested to send through the link of:

by writing the subject line on:

Technical offer QR:                                 Financial offer QR

Qatar Chairty shall not consider any bid that arrives after the deadline for submission. Any bid received by Council after the deadline for submission of bids shall be declared late and rejected.

QC reserves the right to enquire compliance with additional conditions/requirements as and when issuing the final solicitation documents. Submitting a reply to a call for EOI does not automatically guarantee receipt of the solicitation documents when issued.

Supplier Authorized Signatory

Name:

Title:

Signature: